



**Firmware Update**

From time to time Vescent will issue new releases of firmware for the SLICE products. Below are instructions on upgrading the SLICE-QT firmware. Please follow these instructions carefully.

|  |  |
| --- | --- |
|  | For SLICE-QT firmware versions 1.38 and lower, updating the firmware will reset all stored parameters (PID loop parameters, etc.). Before beginning a firmware update, manually record all settings for all channels. Versions 1.50 and higher will automatically store this information. |

1. Create the following folder on your PC C:\Vescent\SLICE-QT\SLICE-QT-0001
2. Copy the latest Upgrade Package from the ftp website (email sales [at] vescent [dot] com for latest firmware and GUI revisions).
3. Unzip the Upgrade Package into the folder created in step 1. There will be two files in the zip file.
   1. One will be named SystemController-x-x.dfu where x-x will be the version number of the firmware.  
      **NOTE**: This file may be renamed. However, the name must start with SystemController.
   2. The other file will be named 90-00001-upgrade-x-x.hex where x-x will be the version number of the firmware.  
      **NOTE**: This file may be renamed. However, the name must start with 90-00001.
4. Determine the COM port number for the SLICE unit’s USB connection   
   (You can skip this step if you already know which COM port is assigned to your SLICE unit.)
   1. If the SLICE unit’s USB cable is connected to the PC, **unplug it**.
   2. Open a Windows Command Prompt console
      1. Press the Windows Start Button
      2. Type **cmd.exe** in the box that says “*Search programs and files”*
      3. Press the <Enter> key
      4. At the command prompt type “**powershell**” and press the <Enter> key   
         You should see a window like the one shown below with **PS** preceding the command prompt.  
         
      5. Highlight the following line and use <CTRL> + <C> to copy it to the Windows clipboard  
         [System.IO.Ports.SerialPort]::getportnames()
      6. Paste it into the Command Prompt window by Right Clicking anywhere in the window and selecting **Paste**Then press <Enter>  
         
      7. You should see a list of all your machine’s COM ports as shown below  
         
      8. Connect the USB cable from the PC to the SLICE and power on the SLICE
      9. Press the <Up-Arrow> key which will repeat the command from step vii above and press <Enter>. The COM port assigned to the SLICE should now appear in the list.   
         NOTE: COM11 is the SLICE COM port in the example below but it may be a different number on your machine.  
         
5. Connect a USB cable (type B) from the back of the SLICE unit to your PC
6. Do this step only the first time you use a new PC to do the upgrade:
   1. Power OFF the SLICE-QT
   2. Depress and hold the blue button within the access hole on the lower left front side of the SLICE case, located just behind the red valance.  
      
   3. While holding the blue button, power on the SLICE. You should see a blank white screen on the display.
   4. Wait for Windows to acknowledge the device found and install the drivers.
   5. Power Off the SLICE-QT
7. Power On the SLICE-QT
8. Copy the latest version of PC Host GUI from the FTP site (see above)
   1. Copy SLICE-QT-Windows.exe to your local PC
   2. Double click on SLICE-QT-Windows.exe to launch the application. You should see the window in [figure 1](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_home).




Fig. 1: Home Screen

1. Click on the ‘OFF’ button below COM Port. You should see a menu similar to the one shown [figure 3](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_com1). NOTE:The COM port numbers may differ from those shown below.




Fig. 3: COM port selection

1. Click on the COM port in the list that is assigned to your SLICE unit. (COM11 in this example). The window in [figure 4](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_com2) should appear if the connection is successful.

https://www.vescent.com/manuals/ Printed on 2019/01/30 01:01




Fig. 4: Ready to upgrade

1. Click on the Upgrade button below the version numbers to see [figure 5](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_upgrade). Note that multiple SLICE-QT-000X folders can be created on your hard drive in the C:Vescent\SLICE-QT folder. The upgrader picks the one with the highest number (SLICE-QT-0002 in the example shown below).




Fig. 5

1. Click Continue to reach the screen shown in [figure 6](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_blue_butt).
   1. Follow the instructions on this screen.

NOTE: The blue button is behind the small hole on the lower left side of the SLICE just behind the red valance.

Last update: 2019/01/16 15:14 slice:qt:firmwarehttps://www.vescent.com/manuals/doku.php?id=slice:qt:firmware




Fig. 6

1. Click Continue to start the upgrade process.

NOTE: During the upgrade process, the SLICE will reset itself after the System Controller code is loaded.

DO NOT attempt to power off or use the SLICE box yet. The QT firmware upgrade still needs to finish.

1. During the second part of the update (QT firmware portion), the message in [figure 7](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_upgr_touch) will be displayed on the SLICE touch screen.




Fig. 7

1. When the update is complete, the SLICE unit will reset and the home screen will be displayed on the touch screen.

The PC Vescent SLICE-QT application should display the message in [figure 8](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware#img_complete).

https://www.vescent.com/manuals/ Printed on 2019/01/30 01:01




Fig. 8

1. Verify the firmware versions on the SLICE Settings screen.
2. Close GUI and restart after completion.



Please follow firmware update instructions

carefully.

During the upgrade, do not turn off,

disconnect, or unplug the SLICE.

From:

https://www.vescent.com/manuals

[/](https://www.vescent.com/manuals/)

-

**Product Manuals**

Permanent link:

**https://www.vescent.com/manuals/doku.php?id=slice:qt:firmwar**

[**e**](https://www.vescent.com/manuals/doku.php?id=slice:qt:firmware)

Last update:

**2020/02/28 17:23**

